## Slide 1: TXCH Global HOPE
Pharmacy Cleaning and Maintenance

## Slide 2: Objectives
By the end of this presentation, the participant should be able to:
  Determine the frequency by which pharmacy areas should be cleaned
  Describe the steps necessary to clean the biosafety cabinet
  Summarize the risks of inadequate cleaning and maintenance

## Slide 3: Pharmacy Area Cleaning: Daily
The following tasks should be completed at least once a day:
  Ensure floors and counters are clean and tidy
  Remove trash and waste bins (including hazardous and sharps containers) to appropriate disposal location
  Clean Biosafety cabinet at start and end of a shift

## Slide 4: Pharmacy Area Cleaning: Weekly
The pharmacy workroom, ante room, and clean room should be deep/“terminal” cleaned at least once a week
Ensure all areas are cleaned:
  Ceilings
  Walls
  Sinks
  Counters
  Tables
  Sides and legs of the biosafety cabinet
  Chairs/stools
  Air vents/returns
  Floor

## Slide 5: Pharmacy Area Cleaning: Monthly
The following tasks should be completed at least once monthly:
  Remove all supplies/medications from storage locations (e.g., shelves, refrigerators)
  Wipe down all bins, cabinets, shelves
  Wipe all medication preparation areas
  Replace all supplies/medications

## Slide 6: Biosafety Cabinet Maintenance
Maintenance
All biosafety cabinets (BSCs) should be inspected a minimum of every 6 months
BSC inspection should be completed by a NSF-certified external vendor
Cleaning & Decontamination
Clean the BSC at the beginning and end of each shift, when a spill occurs, or if the blower was turned off
Decontaminate the BSC at least once a month and
  When a spill enters the air intake vents
  Before moving or servicing the cabinet
NSF = National Sanitation Foundation; www.nsf.org

## Slide 7: Biosafety Cabinet Cleaning
Alternate cleaning materials/chemicals each week, per table below
Note: Cleaning procedure steps are on the following slide
*Sterile gauze should be 4 inch x 4 inch in size, and 4-ply
  Do not use other materials (wool, sponges, etc.) as they will leave non-sterile particulate matter
Week | First Step | Second Step
1 | Spray 10% bleach solution and allow to remain on the surface for 2 minutes | Wipe (per cleaning procedure) with sterile gauze* soaked with sterile 70% isopropyl alcohol to remove all residue
2 | Use Envirocleanse spray (per cleaning procedure) and allow to remain on the surface for 2 minutes | Wipe (per cleaning procedure) with sterile gauze* soaked with sterile 70% isopropyl alcohol to remove all residue

## Slide 8: Biosafety Cabinet Cleaning Procedure
Back wall
Sides
Glass shield
Equipment
Work surface
Clean the horizontal pole (if present) using a slow, continuous movement in one direction
Start at the top back corner of each side, using continuous overlapping strokes in a horizontal direction until the bottom of the side is reached
Reach as far as possible up the side of glass shield and use continuous overlapping strokes in the horizontal direction until the bottom of the glass shield is reached
Equipment in the hood must be cleaned and should be moved to allow for cleaning of the work surface underneath the equipment
Start the back corner, use continuous overlapping strokes, moving in the horizontal direction until the front of the hood is reached
Use continuous overlapping strokes starting at the top back corner and moving horizontally back and forth
Pole
These steps are for a vertical laminar airflow cabinet; different procedures are required for a horizontal laminar airflow cabinet

## Slide 9: Biosafety Cabinet Decontamination
Assemble the following supplies
  Screwdriver
  Chemo gloves (2 pair)*
  Chemo gown (1)*
  Dust and mist respirator mask (1)*
  Disposable safety glasses (1)*
  Chemo waste bag (2)*
  Spill towels (3)*
  Chemo spill pads (2)*
  Shoe covers (1 pair)*
  Sterile water for irrigation (2 x 1 liter bottles)
  Sterile gauze pads
  Bleach solution or Envirocleanse (depending on which product is in use for cleaning that week)
*These items may be present in a Chemo Spill Kit

## Slide 10: Biosafety Cabinet Decontamination
Follow the correct procedures for Donning Personal Protective Equipment and Hand Scrubbing prior to beginning the Decontamination
Remove screws from the work tray, place them in a safe location
As you clean, all used materials (gauze, gloves, etc.) should be placed in a ziplock bag, then lastly in the chemo waste bag for disposal
Take care to not wet the HEPA filter

## Slide 11: Biosafety Cabinet Decontamination
Clean the BSC in the following order:
  Top filter grill
  Back wall
  IV bar/pole
  Side walls
  Work tray (including the back side/underneath)
  Inside/bottom cabinet panels
  Glass shield (inside)
  Glass shield (outside) plus other outer surfaces
Use bleach (or Envirocleanse) to wash and scrub all areas
  After washing, rinse/wipe with sterile water and gauze
  Use new sterile gauze pads frequently
  Dispose of gauze in ziplock bag

## Slide 12: Cleaning & Maintenance Records
Complete the appropriate/corresponding Pharmacy Records Forms when any of the following are completed:
  Cleaning
    Daily
    Weekly
    Monthly
  Decontamination
  Maintenance/certification
  Aseptic technique validation

## Slide 13: Cleaning & Maintenance Summary
Frequency | Task
Each shift | Clean the biosafety cabinet
Daily | Ensure pharmacy areas are clean and tidy; remove trash and waste
Weekly | Terminal cleaning of pharmacy areas
Every 6 months | BSC inspection and certification by qualified vendor
Annually | Aseptic technique re-validation
If any signs of failure to adhere to aseptic technique, or if the BSC is not re-certified, then additional steps are required to ensure re-validation of technique and re-certification of the BSC

## Slide 14: What’s Next?
Watch video on cleaning the biosafety cabinet
Complete practice questions
Review answer fields

## Slide 15
Global HOPE Pharmacy Education